| ###### Institution: | ###### Medstar Georgetown University Hospital - Clinton |
| --- | --- |
| ###### Meeting Date: | ###### November 03, 2025 |
| Meeting Time | ###### 11:30 AM Eastern Time |
| ###### Meeting Type: | ###### Virtual Platform Teleconference (Remote) ###### Open to the Public |
| Members in Attendance: | Member; Voting; Member Type. Noriea, Nicholas; Yes; Chair: Biosafety Expert/HGT Expert. Rastein, Daniel; Yes; Core Member: Biosafety Expert/HGT Expert. Ellis, Robert; Yes; Core Member: Biosafety Expert/HGT Expert. Browne, Angella; Yes; Local Unaffiliated Member. Bland, Monica; No; Site Contact. |
| Invited Members Not in Attendance: | Member; Voting; Member Type. Morland, Melissa; Yes; Local Unaffiliated Member. |
| #### Guests: | #### Akinsanya, Jemima; Yashinski, Cynthia; Philip, Sherin; Cavenaugh, Carlette; Elgendi, Sarah |
| #### Staff: | #### Sreedharan, Aswathy; Payne, Kaylie |

Call to Order: The IBC Chair called the meeting to order at 11:31 AM Eastern Time A quorum was present as defined in the Sabai IBC Charter.

Conflicts of Interest: The IBC Chair reminded all members present to identify any conflicts of interest (COI).

Public Comments: No public comments were made prior to or at the meeting.

Review of Prior Business: None

Previous Meeting Minutes: None

New Business:

| PI: | Akinsanya, Jemima DO |
| --- | --- |
| Sponsor: | ModernaTX, Inc. |
| Protocol: | mRNA-1195-P201 A Phase 2, Randomized, Observer-Blind, Placebo-Controlled, Dose-Ranging Study of mRNA-1195 Intramuscular Injection in Participants 18 to ≤55 Years of Age With Multiple Sclerosis |
| Review Type: | Initial Review |
| NIH Guidelines Section: | III-C-1 |

Trial Summary: mRNA-1195-P201 is a randomized, observer-blind, placebo-controlled Phase II

study sponsored by ModernaTX, Inc. and designed to assess the safety and reactogenicity of

mRNA-1195 in adult participants with clinically isolated syndrome (CIS), radiologically isolated

syndrome (RIS), or multiple sclerosis (MS). mRNA-1195 is intended for use as an Epstein-Barr

virus (EBV) vaccine and consists of lipid nanoparticle (LNP) formulated messenger RNAs

(mRNAs) encoding EBV antigens. The investigational product (IP) is administered by IM injection.

Biosafety Containment Level (BSL): The study agent mRNA-1195 consists of synthetic

mRNAs that do not encode any known hazardous transgenes or toxins, therefore BSL-1

containment is the minimum recommended biocontainment level.

Risk Assessment and Discussion:

- The Committee reviewed the clinical trial Sponsor’s study documents and the Sabai-generated comprehensive study-specific Risk Assessment which collectively provided a thorough description of the recombinant or synthetic nucleic acid molecules (investigational product/s) and the proposed clinical research activities involving the IP.

  - In summary, the primary risks in this clinical trial include potential occupational exposure from accidental spills or splashes of the IP during preparation and/or administration procedures and needlesticks due to the use of needles during preparation and/or administration. These potential risks are mitigated through a combination of relevant staff training, safe clinical practices (including Standard Precautions and sharps safety) and use of appropriate PPE (as prescribed in the Risk Assessment and documented in the IBC submission package)

  - The Site confirmed that only study personnel who have been educated on the potential biohazards and the precautions to be taken when working with the IP will handle the IP or any materials contaminated by the IP.

  - The Site confirmed that study personnel are sufficiently trained in the practices and techniques required to safely work with the IP.

  - The Site confirmed that staff members receive Bloodborne Pathogens training.

  - Occupational Health Recommendations: None

  - The Committee had no additional significant comments or recommendations regarding the description of the potential risks and occupational exposure hazards associated with handling the IP in this clinical trial, or the proposed mitigation strategies, as detailed in the Risk Assessment.

- The Committee reviewed the Site’s facility details, relevant study-specific procedures and practices, the PI’s credentials and other applicable information provided by the Site for the purposes of the IBC review.

  - The Site verified that the information provided by the Chair was accurate.

  - The Committee noted some equipment located inside the BSC and recommended that the Site ensure that the air vent grills of the BSC are not blocked. Additionally, the Committee reminded the Site to ensure that any equipment that will stay permanently inside the cabinet is left inside at the time of vendor certification.

  - In response to a question from the Committee, the Site confirmed that their standard practice is to not place a sharps container inside the BSC, and that the Pharmacy plans on using CSTDs for preparation. The Committee suggested that the Site place a sharps container inside the BSC if the Sponsor requires the use of standard sharps. The Site had no concerns.

  - The Committee recommended that the Site consider placing a physical barrier between the BSC and the supply rack to reduce the potential for contamination of supplies.

  - The Site confirmed the participant would be dosed on the exam table. The Site also confirmed the chairs in the room may have absorbable backing. The Committee requested that the Site either remove or cover any chairs with absorbent fabric, or alternately, given the small volume of agent, ensure that the agent is handled at least 6 feet away. The Site noted they would remove the chairs, and the Committee had no questions.

  - In response to a question from the Committee, the Site confirmed they would affix a larger eyewash placard to the wall behind the eyewash in the dosing room, similar to the one in the Pharmacy.

  - The Site confirmed that there is a biohazard label on non-sharps dosing waste container.

  - The Site confirmed once the storage unit is received, they will send a photo of the unit with biohazard sign affixed.

Motion: A motion of Full Approval for the study at BSL-1 plus Standard Precautions was passed by unanimous vote. There were no votes against and no abstentions.

- Contingencies stated by the Committee: None

- Stipulations stated by the Committee: None

Review of Incidents: Nothing to report.

IBC Training: Nothing to report.

Reminder of IBC Approval Requirements.

Adjournment: The IBC Chair adjourned the meeting at 12:31 PM Eastern Time

Post-Meeting Pre-Approval Note: None